## Slide 1: SPECT
Dr. Julia Mannheim
julia.mannheim@med.uni-tuebingen.de

## Slide 2: Overview SPECT
general principle
collimation
single vs. multi-pinhole collimation
data acquisition and signal processing
SPECT in comparison to PET

## Slide 3: SPECT principle
Courtesy: Dr. Simon Cherry, UC Davis
Forms of Nuclear Imaging
SPECT
PET

## Slide 4: Gamma Radiation
a nucleus in an excited state decays to ground state
electron capture: the nucleus consists of too many protons
and cannot emit a positron  captures an electron
e.g. 111In 111Cd + γ
beta decay: emission of γ
e.g. 67Cu 67Zn + e- + γ
Gamma-Ray
e.g. 99mTc 99Tc + γ

## Slide 5: Projection Imaging
Courtesy: Dr. Simon Cherry, UC Davis
Mouse injected with radiotracer
Position-sensitive scintillation detector
Mouse injected with radiotracer
Position-sensitive scintillation detector
Projection image
Position-sensitive scintillation detector
Projection image
Mouse injected with radiotracer
Position-sensitive scintillation detector
Projection image

## Slide 6: Projection Imaging
Courtesy: Dr. Simon Cherry, UC Davis
(Slit collimator)
Mouse injected with radiotracer
Position-sensitive scintillation detector with slit collimator
Projection image

## Slide 7: Collimation
Collimation = direction of motion become aligned
 production of parallel γ -rays
Material of collimator:
  Tungsten, lead

## Slide 8: Pinhole collimation
Parallel-hole collimator
One-to-one Projection
Large FOV
Pinhole collimator
Magnification of object
Smaller FOV
Increased resolution
Collimator
Detector
Object
Magnification
Object

## Slide 9: Pinhole collimation
Magnification
Object
use pinhole to obtain projection
magnification possible due to distance variation
Object

## Slide 10
Courtesy: Siemens
Pinhole collimation
Notes: 10

## Slide 11
Courtesy: Siemens
Pinhole collimation
Notes: 11

## Slide 12: Multipinhole collimation
Object
higher sensitivity
faster acquisition
lower dose
more efficient use of detector
no reduction of resolution
larger FOV

## Slide 13: Inveon Collimators
Rat Mouse
Rat Mouse

## Slide 14: Inveon Collimators
radius of rotation (ROR)
Courtesy: Siemens

## Slide 15: Data Acquisition
Rotate imaging system to collect projection views around the object
Use reconstruction algorithms to determine 3D distribution of radiotracer
Multiple pinholes often used to improve sensitivity

## Slide 16: Data Acquisition
projection 1
acquisition
time
projection 2
.
.
.
.
.
projection xy

## Slide 17
Electronic
Photomultiplier
absorption of γ-rays in scintillation crystals → light release
transformation of light into electrons at photocathode of photomultiplier
amplification of electronic signal in photomultiplier
position-encoding electronics
2.2 mm x 2.2 mm crystals scanning area 150 mm x 150 mm
3x3 array position sensitive pmt
Detector
Courtesy: Siemens

## Slide 18: Data Acquisition & Processing
projection 1
projection 2
projection XY
projection 1
projection 2
.
.
.
.
projection xy
List mode data
Time
Histogram
Projections
Image
Considerations:
Different pinhole configurations allow for choice of different sensitivity and image resolution
More pinholes, lower resolution  higher sensitivity
Recon

## Slide 19: Iterative Reconstruction
Image courtesy of S.R. Cherry, University of California

## Slide 20: Iterative Reconstruction
Image courtesy of S.R. Cherry, University of California

## Slide 21: Energy discrimination
140
200
300
0
Energy
Window
Different SPECT Isotopes can have one or more discrete lines of energy emission (~30-300 keV suitable for the Inveon system)
Post acquisition energy discrimination is performed
Detector needs to be calibrated for each isotope!!
Potential use of multiple isotopes

## Slide 22: SPECT isotopes
isotope | photopeak [keV] | suggested energy window [keV] | half life
57Co | 122 | 110 – 134 | 271.74 d
67Ga | 93 | 84 – 102 | 3.3 d
123I | 159 | 143 – 175 | 13.2 h
125I | 35 | 25 – 45 | 59.4 d
111In | 171 | 154 – 188 | 2.8 d
177Lu | 208 | 187 – 229 | 6.73 d
99mTc | 140 | 126 – 154 | 6 h
201Tl | 70 | 63 - 77 | 12.23 d

## Slide 23
Multi-isotope SPECT
energy
window 1
126-154 keV
energy
window 2
143-175 keV
Adapted from Takeuchi, W. et al. EJNMMI Phys. 2016;3:10
Multi-isotope imaging, e.g., 99mTc (peak 140 keV) & 123I (peak 159 keV)

## Slide 24: Multi-isotope SPECT
Courtesy: Siemens

## Slide 25: PET vs. SPECT
 | PET | SPECT
spatial resolution | crystal thickness scintillation material depth of interaction reconstruction parameters … | pinhole diameter scintillation material radius of rotation reconstruction parameters …
sensitivity | crystal length axial FOV size … | pinhole diameter number of pinholes radius of rotation …

## Slide 26
Spatial resolution
  PET: 1.5 mm CFOV
  SPECT: 0.8 mm, 25 mm ROR, 1MHR-0.5
  1.2 mm, 25 mm ROR, 1MGP-1.0
        1.3 mm, 30 mm ROR, 5MWB-1.0
  depending on pinhole choice  PET worse spatial resolution compared to SPECT
Sensitivity
  PET: 7.2% at CFOV (energy window: 350-650 keV,
        timing window: 3.432 ns)
  SPECT: 0.0107%, 25 mm ROR, 1MGP-1.0
        0.0286%, 30 mm ROR, 5MWB-1.0
  PET better sensitivity than SPECT
values represent the Inveon system
PET vs. SPECT