## Slide 1: Image Reconstruction
Jonathan Disselhorst, PhD
Jonathan.Disselhorst@med.uni-tuebingen.de

## Slide 2: Downloads
http://jonathan.disselhorst.eu/imageanalysis
(case sensitive)

## Slide 3: Overview
Forward and back projection
Reconstruction types
Filtered Back Projection
  Filtering
Ordered Subsets Expectation Maximization
  Iterations
  Subsets
MATLAB

## Slide 4: Forward Projection

## Slide 5: Forward Projection
1 projection,
1 angle
Sinogram of
1 slice
over 180o

## Slide 6: Forward Projection

## Slide 7: Forward Projection

## Slide 8: Forward Projection
Sinogram
Image

## Slide 9: Back Projection
A+B+
A+D+
A+C+
A
A+B+
B+
B+D+
B+C
C+D+
A+D+
B+D+
D
C+D+
C+
A+C+
B+C

## Slide 10: Back Projection
A+B+
A+D+
A+C+
A
A+B+
B+
B+D+
B+C
C+D+
A+D+
B+D+
D
C+D+
C+
A+C+
B+C
Subtract projection sum (A+B+C+D) from all pixels
Divide by number of projections − 1
A
B
D
C

## Slide 11: Back Projection
Sinogram
Reconstructed Image
(unfiltered back projection)

## Slide 12: Back Projection
With normal back projection, the image looks very blurry!
Blurred with an (1/r) function.

## Slide 13: 1/r
Image of a point (source)
Profile
original
Unfiltered backprojection

## Slide 14: Back Projection
With normal back projection, the image looks very blurry!
Blurred with an (1/r) function.
Could be corrected with a deconvolution, but we don’t want to do this…

## Slide 15: Reconstruction types
Image Reconstruction
Analytical
e.g., FBP
Iterative
Statistical
e.g., OSEM
Algebraic
e.g., ART

## Slide 16: Fourier Transform
 Discrete Fourier transform 
 Inverse discrete Fourier transform 
Discrete form of the Fourier transform (DFT)
Often implemented as the fast Fourier transform (FFT)
High frequency
Low frequency

## Slide 17: Central Slice Theorem
Cherry et al. Physics in Nuclear Medicine
Can be used for reconstruction.
Reconstruct image by applying the inverse Fourier transform
However, converting the polar frequency space to Cartesian coordinates is slow and can lead to artifacts when inter-polation is not done properly.
Notes: The central slice theorem is an important concept in image reconstruction.
Projection p(r,ϕ) is a one-dimensional projection of the 2D image f(x,y) at projection angle ϕ.
Fourier spectrum F(kx,ky) is the Fourier transform of image f(x,y).
The theorem states that the 1D Fourier transform of p [called P(kr,ϕ)], is equal to a line through F at the same angle ϕ.

## Slide 18: Filtered Back Projection
Take 1D FFT of each projection
Perform filtering in frequency domain (does not require interpolation or coordinate conversion: therefore it is accurate and fast)
Take 1D inverse FFT
Standard back projection

## Slide 19: Filters
What would be a good filter?
High frequency
Low frequency
H
k
Ramp filter (a.k.a. Ram-Lak)

## Slide 20: Filters
Back projection
Filtered back projection
(Ramp filter)

## Slide 21: Filters
Image courtesy of Chih-Chieh Liu
Low frequencies:
Always filtered out

## Slide 22: FBP vs. Reality
Imprecise detectors
Imperfect geometry
…
Number of projections
Noise

## Slide 23: Number of Projections
Point source measurement reconstructed from 1 projection.
Point source measurement reconstructed from 2 projections.
Point source measurement reconstructed from 3 projections.
Point source measurement reconstructed from 6 projections.
Point source measurement reconstructed from 12 projections.
Point source measurement reconstructed from 24 projections.
Point source measurement reconstructed from 48 projections.
Point source measurement reconstructed from 96 projections.
Filtered
Unfiltered

## Slide 24: Noise
Image courtesy of
Milan Zvolský (DESY)
Number of events follows a Poisson distribution
Signal-to-noise ratio
Analytical reconstruction works exactly when

## Slide 25: Noise
Sinogram
Ramp filter
Hann filter

## Slide 26: Noise
Shepp-Logan filters with different cut-off frequencies.
Cherry et al.

## Slide 27: Artifacts
Effect of linear sampling of the profile on image quality.
Undersampling leads to blurring and image artifacts.
Cherry et al.

## Slide 28: Artifacts
Cherry et al.
Effect of angular sampling on image quality.
Undersampling leads to image artifacts.

## Slide 29: Artifacts
Effect of angular sampling range on image quality.
Sampling over an interval of less than 180 degrees leads to image artifacts.
Cherry et al.

## Slide 30: Artifacts
Cherry et al.
Projections do not cover the entire object.

## Slide 31: Artifacts
Cherry et al.
Missing projection element, e.g., as a result from a broken detector.

## Slide 32: Iterative Reconstruction
Cherry et al. Physics in Nuclear Medicine
Cost function
Update function
The measurement,
Forward projection

## Slide 33: Iterative Reconstruction
Cherry et al. Physics in Nuclear Medicine
Aspects that may be considered during iterative reconstruction:
Statistical noise
Prior information (about the system, the object, …)
Non-negativity constraint

## Slide 34: Iterative Reconstruction
Cherry et al. Physics in Nuclear Medicine
Major advantages:
Less sensitive to noise
Potentially more accurate
Major disadvantages
Computationally intensive

## Slide 35: Iterative Reconstruction
Cherry et al. Physics in Nuclear Medicine
Iteration 1
Iteration 3
Iteration 5
Iteration 10
Iteration 20
Iteration 30

## Slide 36: Iterative Reconstruction
Components in iterative reconstruction:
Image model
System model
Data model
Model of physical effects that govern the system, e.g.,
Geometry
Depth of interation
Scattering in detector
Dead time
Randoms
Photon acollinearity
Positron range
Data model:
Describes the statistical relationship between the measured data and the expected value of the measurement.
Photon detection is Poisson distributed* 
*no longer after correction for randoms, scatter, attenuation, etc…

## Slide 37: Iterative Reconstruction
Components in iterative reconstruction:
Image model
System model
Data model
Cost function
Algorithm
Most commonly used for iterative reconstruction: Maximum Likelihood.
In general: the likelihood that observed data matches with a certain model.
An algorithm is required to optimize the cost function, to find the best image estimate. Commonly used: Expectation Maximization (EM)

## Slide 38: Iterative Reconstruction, Noise
Ideal
Noisy
Smoothing
Wernick and Aarsvold. Emission Tomography: The fundamentals of PET and SPECT.

## Slide 39: Ordered Subsets EM (OSEM)
Ordered Subsets Expectation Maximization (OSEM)*
Faster than MLEM
Uses subsets of the entire data for each image update.
Each iteration has sub-iterations for each subset
The back projection only uses projections from one subset
* Hudson and Larkin. IEEE Trans Med Imaging. 1994

## Slide 40: OSEM
Rule of thumb: Speed increase of OSEM is roughly the number of subsets.
Caution: comes at the expense of higher image variance
Alessio and Kinahan, PET Image Reconstruction. 2006
OSEM with one subset = MLEM

## Slide 41: FBP vs. OSEM
Alessio and Kinahan, PET Image Reconstruction. 2006
FBP vs. OSEM
FBP vs. OSEM

## Slide 42: Fan-Beam & Cone-Beam
Fan-beam
Cone-beam
Cherry et al.
Feldkamp algorithm

## Slide 43: 3D – PET
Cherry et al.
Single Slice Rebinning – Fourier Rebinning – 3D reconstruction

## Slide 44: MATLAB
IM = phantom;
sino = radon(IM);
FBP = iradon(sino);
CT = fanbeam(IM,183);
CTrecon = ifanbeam(CT,183);
sinon = imnoise(sino/1E12,'poisson');
FBPn = iradon(sinoN);
183 is the distance from CT source to the center of rotation. Can be changed.
1E12 is a scale factor to adjust the noise, increase to increase the noise. Also ‘Poisson’ can be changed to change noise distribution.

## Slide 45: Questions!
http://jonathan.disselhorst.eu/imageanalysis
(case sensitive)
Downloads

## Slide 46: Books